Contact: sales@cepoint.com 603.883.7979. www.cepoint.com

DAS-6063T rugged, portable, multifunctional data acquisition system for temperature strain, accelerometer, transducer, vibration testing and measurement.

DAS-6063T  TransPAC  is a hardened, ruggedized  portable  16-Channels Analog/Digital, multifunctional data acquisition system, for testing , measurement, acquisition or data logging. Whether the applications is Temperature  transducer, strain, accelerometer or vibration testing or measurements, DAS-6063T is built o handle the though demands. The system  hardware is comprised of up to optional 80 I/O plug-and-play PCI board, 3 ft. cable, and screw terminal/signal conditioning pod. Cold-junction compensation is built-in to the

DAS-6063T, which provides precise thermocouple

measurements, with channel-to-channel variation and

noise less than 0.2°C, and accuracy better than 1.0°C. A

separate low-noise amplifier per TC channel allows for

these exact measurements. DAS-6063T is available with

either 7, 14 or 28 or 32 thermocouple input channels and an

8-channels RTD.

In addition to Analog and thermocouple inputs, DAS-6063T also provides

7 or 14 differential voltage input channels that offer 16-bit resolution

and 13 programmable full-scale input ranges from

±156 mV to ±10V. per port .The 4 or 8 frequency input channels can

measure frequency or count pulses up to 10 MHz. Two

timer output channels are capable of generating square

wave outputs up to 1 MHz. 24 or 48 digital I/O lines are

also included, which can be programmed in 8-bit groups as

either inputs or outputs.

All inputs, including thermocouple inputs, voltage, frequency, and digital, can be synchronously scanned up to 200K readings/s, which provides accurate time correlation

between all inputs. The 16-bit/100-kHz analog outputs can be used to generate waveforms with a virtually infinite waveform output length from the system memory.

All inputs and outputs are accessed via convenient screw

terminals located on the signal conditioning pod. The cover

of the signal conditioning pod is easily removed to gain

access to the screw terminals for signal connections.

Screw-terminal access to all voltage inputs, digital I/O, frequency I/O,

and analog output

DAS-6063Tincludes a full suite of third party software

including DaqView™ Out-of-the-Box™ software, plus drivers for DASYLab®, LabVIEW®,

TestPoint®, MATLAB, and Windows®-based

languages including Visual Basic®, and C++

ADaqTemp Series Configurations*

Applications:

Production test

Vibration analysis

Stimulus/response analysis

Process monitoring and control

Temperature

Sound analysis

Automotive engine testing

FEATURES OVER VIEW:

- Rugged aluminum alloy hardened industrial chassis

- Up to 16 or 32-Channel, 200 KS/s A/D

- A/D 16-bit resolution

- 8- Digital I/O bits & 2 counter/timer

- Built with powerful Intel Pentium III- 1GHz  processor

- 256 MB internal DDRAM memory

- 80GB internal storage shock mount hard disc drive

- (Removable or swappable storage option)

- 3.5” floppy and 24 x CD-ROM

- Flash Disk option

- 12- 15” TFT LCD display 1024 x768 with analog video converter

- Display of multiple windows or reports

- Up to 4 simultaneous window display view of tests or reports

- Real-Time data acquisition or logging.

- Data transfer or transport via removable media, TCP/IP or RS-232/422.

- Reports or results in ASCII protocol, Excel spreadsheet

Series

Specifications & Ordering Information

IO, Inc. • 25971 Cannon Road • Cleveland, OH 44146 • 440.439.4091 • Fax 440.439.4093 • s@iotech.com • www.iotech.com

The DAS-6063T provides 7,14, 28 or 32 differential channels

of TC/mV inputs (depending on factory configuration), and

8 RTD channels configured in 37-pin ports: P1, P2, P3. to

physically connect the thermocouples or RTD’s via DBK screw

box terminals that are signal conditioned.

System Connector: All DBK options have a DB37 male,

which mates with P1, P2, P3 on the Das-6063T system, or

other DBK options

TC/mV Connector

DBK81: Board-mounted screw terminals

DBK82: Board-mounted screw terminals

DBK83: External pod-mounted screw terminals

DBK84: Mini-TC connectors

Functions: TC types J, K, S, T, E, B, R, N; x100 (voltage)

Inputs

DBK81: 7 differential TC/mV inputs

DBK82: 14 differential TC/mV inputs

DBK83: 14 differential TC/mV inputs

DBK84: 14 differential TC/mV inputs

Input Voltage Range: 100 mV**

Input Impedance: 40M Ohm (differential); 20M Ohm

(single-ended)

Input Bandwidth: 4 Hz

Input Bias Current: 10 nA typ

CMRR: 100dB typ

Maximum Working Voltage (signal + common

mode): 10V

Over-Voltage Protection: 40V

Power Requirements

DBK81: 35 mA max from 15V;

2 mA max from +5V

DBK82: 60 mA max from 15V;

2 mA max from +5V

DBK83: 60 mA max from 15V;

2 mA max from +5V

DBK84: 60 mA max from 15V;

2 mA max from +5V

Operating Temperature: 0ªC to 50ªC

Voltage Accuracy: (0.2% of rdg +50 V)

TC Accuracy: Valid for one year, 18¢ª to 28¢ªC†,

see table below

Minimum Resolution: 0.1ªC for all TC types

TC Accuracy table:

Type	 Min	 Max 	-100	 0 	100	 300	 500	 700 	900	 1100 	1400

J	 -200	760	 0.6	 0.5	 0.5	 0.6	 0.7	 0.7	 —	 — 	—

K	 -200 	1200	 0.7	 0.6	 0.6 	0.7	 0.8	 0.8	 0.9	 1.0	 —

T	 -200 	400	 0.7	 0.6	 0.6	 0.6	 —	 —	 —	 —	 —

E	 -270 	650 	0.6	 0.5	 0.5	 0.5	 0.6	 — 	— 	—	 —

S	 -50	 1768	 — 	2.3 	1.8	 1.5	 1.5	 1.4	 1.5	 1.6 	1.6

R 	-50 	1768	 — 	2.3 	1.6	 1.5 	1.4	 1.4	 1.3	 1.4	 1.5

B	 50	 1780	 — 	—	 —	 3.7	 2.4	 2.1	 1.8	 1.7	 1.5

N28	 -270 	400	 0.9 	0.7	 0.7	 0.7	 —	 —	 —	 — 	—

N14	 0 	1300	 — 	0.7	 0.7	 0.7 	0.8	 0.8 	0.9 	1.0	 —

General

Power:  110-250ac auto-switching, 250 watts. 24 or 48 DC option available

Operating Temperature: 0° to +60°C

Relative Humidity: 0 to 95%, non-condensing

System Dimensions:

Dimensions: 17.25” W x 13.38” H x 9.25” D

Weight: approx. 35lbs

Signal Conditioning/Screw Terminal Pod

Dimensions: 86 mm W x 240 mm D x 42 mm H

(3.4” x 9.44” x 1.65”)

A/D Specifications

Type: Successive approximation

Resolution: 16 bit

Conversion Time: 5 µs

Maximum Sample Rate: 200 kHz

Nonlinearity (Integral): ±1 LSB

Nonlinearity (Differential): No missing codes

Voltage Inputs

Channels: 7 differential inputs, programmable on a

per-channel basis for unipolar or bipolar

Temperature Coefficient: ±(10 ppm +0.3 LSB)/°C

Input Impedance: 20M Ohm (differential)

Bias Current: <1 nA (0° to 35°C)

Common Mode Rejection: 86 dB, DC to 60 Hz for

gains < =8; >100 dB for gains > =16

Maximum Input Voltage: ±11V relative to analog

common

Over-Voltage Protection: ±35V

Ranges: Software or sequencer selectable on a perchannel

basis, from 156 mV FS to 10V FS

Voltage Accuracy: One year, 0-35°C (exclusive of

noise); for full-scale input ranges from 2.5V to

10V, accuracy is 0.015% of reading +0.005% of

range. For full-scale input ranges from 312 mV to

1.25V, accuracy is 0.015% of reading +0.008% of

range. For full-scale input of 156 mV, accuracy is

0.02% of reading +0.008% of range.

Thermocouple Inputs

Channels: 7,14, 28 or 32 differential TC inputs

TC types: J, K, S, T, E, B, R, N, or mV input

Input Voltage Range: ±100 mV

Input Impedance: 40M Ohm

Input Bandwidth: 4 Hz

Input Bias Current: 10 nA

CMRR: 100dB typ

Over Voltage Protection: ±40V

Voltage Accuracy: ±(0.2% of rdg + 50 µV)

TC Accuracy: See table above; valid for one year,

18°C to 28°C.

Minimum TC resolution: 0.1°C for all TC types

Input Sequencer

Temperature, voltage, digital and frequency inputs

can be scanned synchronously, based on either an

internal programmable timer, or an external clock

source.

Scan Clock Sources: 2

1. Internal, programmable from 5 µs to 5.96

hours in 5 µs steps

2. External, TTL level input up to 200 kHz max

Programmable Parameters per Scan: Channel

(random order), gain, unipolar/bipolar

Depth: 512 locations

Channel-to-Channel Scan Rate: 5 or 10 µs per channel,

programmable

Triggering

Trigger Sources: 6, individually selectable for starting

and stopping an acquisition. Stop acquisition can

occur on a different channel than start acquisition;

stop acquisition can be triggered via modes 2, 4, 5,

or 6 described below. Pre-trigger is supported with

fixed or variable pre-trigger periods.

1. Single-Channel Analog Hardware Trigger

Any voltage or TC input channel can be software

programmed to be the analog trigger channel

2. Single-Channel Analog Software Trigger

Any voltage or TC input channel can be selected

as the software trigger channel. If the trigger

channel is a thermocouple, then the driver automatically

compensates for the delay required to

calculate the temperature, resulting in a maximum

latency of one scan period.

3. Single-Channel Digital Trigger

A separate digital input is provided for digital

triggering.

4. Digital Pattern Triggering

Pattern triggering is available on the digital input

port. Triggers can be set on equal value, above a

value, below a value, inside a window, or outside a

window. Individual bits can be masked for “don’t

care” condition.

5. Counter/Totalizer Triggering

Counter/totalizer inputs can trigger an acquisition.

User can select to trigger on a frequency or

on total counts that are equal, above, below, or

within/ outside of a window.

6. Software Triggering

Trigger can be initiated under program control.

Analog Output

Up to four analog output channels can be updated

synchronously relative to scanned inputs, and

clocked from either an internal onboard clock, or an

external clock source. Analog outputs can also be

updated asynchronously, independent of any other

scanning in the system. Bus Mastering DMA provides

CPU and system-independent data transfers.

This feature ensures accurate outputs that are irrespective

of other system activities. Streaming from

disk or memory is supported, which allows continuous,

nearly infinite length, waveform outputs (limited

only by available PC system resources).

Channels: 0, 2, or 4

Resolution: 16 bits

Output Voltage Range: ±10V

Output Current: ±10 mA

Offset Error: ±0.0045V max

Gain Error: ±0.01%

Update Rate: 100 kHz max, 1.5 Hz min

(no minimum with external clock)

Settling Time: 10 µs max to 1 LSB for full-scale step

Clock Sources: 4, programmable

1. Onboard D/A clock, independent of scanning

input clock

2. Onboard scanning input clock

3. External D/A input clock, independent of

external scanning input clock we should provide it

4. External scanning input clock

Frequency/Pulse Counters

Counter inputs can be scanned synchronously

along with voltage, TC and digital scanned inputs,

based either on internal programmable timer, or

an external clock source. Bus Mastering DMA provides

CPU and system-independent data transfers,

ensuring data acquisition performance irrespective

of other system activities. Counters can be

configured to clear when read, or to totalize and

clear under program control.

Channels: 4 x 16-bit; cascadable as 2 x 32-bit

Frequency Measurement Rate: 10 MHz max

Input Signal Range: -15V to +15V

Trigger Level: TTL

Minimum Pulse Width: 50 ns high, 50 ns low

Frequency/Pulse Generators

Channels: 2 x 16-bit

RTD Specifications

RTD channels:

- Provides eight RTD channels

- Supports 3- and 4-wire RTDs

- Supports three ranges of RTD resistance

- Features resolution over the full range of -200° to +850°C span

Connector: DB37 male, mates with P2 or P3 pin out on the DAS-6063T;  screw terminals and provided for signal connections

Configurations: 3- or 4-wire

Alpha: 0.00385

Inputs: 8 channels

Temperature Ranges:

100, 500, or 1000 Ohm: -200ª C to +850ªC

RTD Excitation Voltage:

100 Ohm 635 mV

500 Ohm 455 mV

1000 Ohm 235 mV

Accuracy:

100, 500, or 1000 Ohm: 1.5ªC

Resolution:

100, 500, or 1000 Ohm: 0.3ªC

Range and RTD Type Adjustments:

Jumpers on circuit board

Output Waveform: Square wave

Output Rate: 1-MHz base rate divided by 1 to

65535 (programmable)

High-Level Output Voltage: 2.0V min @ -3.75 mA;

3.0V min @ -2.5 mA

Low-Level Output Voltage: 0.4V max @ 2.5 mA

Ordering Information

Description Part No.

All DAS-6063T temperature systems include screw terminal measurement

board, 3 ft. cable, signal conditioning pod, and

software (see below). All programming documentation

is included on the CD ROM supplied with

the system.

7 differential TC inputs* DAS-6063-Temp7

7 differential TC inputs,

plus 2 analog outputs* DAS-6063-Temp 14

14 differential TC inputs* DAS-6063-Temp 28A

28 differential TC inputs,

plus 4 analog outputs* DAS-6063-Temp32

* Additional I/O included with every system includes 7 differential

voltage inputs, 24 digital I/O, 4 counter inputs,

and 2 timer outputs.

Software

Software included with the DAS-6063T consists of DaqView,

Windows drivers for Visual Basic and C++, and drivers for DASYLab®, LabVIEW®, TestPoint®, and MATLAB.

Optional software provides a direct link from DaqView™ and DASYLab

into Excel® spreadsheets DaqView XL and optional LogView software

°C (°C)

Cepoint Networks, LLC  Contact:

sales@cepoint.com. 603.883.7979  www.cepoint.com

Cepoint Networks, LLC.©2003 All rights reserved. Cepoint, DAS-6063T,  is a trademark of Cepoint networks, LLC. DBK series, Daqtemp, Daqview is a trademark of IOtech, Inc. INTEL is a registered trademark of Intel corp. AMD is a registered trademark of AMD corp. All other trademarks are trademarks of their respective owners.